to:		burnside directors

from:		ANSON ANTRIASIAN

subject:	Review of fire assay techniques at north australian labs (nal)

Date:		31/08/2005

cc:

North Australian Laboratories (NAL) was visited Monday, August 29 2005. The entire process of Au fire-assaying was explained and demonstrated from beginning to end courtesy of Ray Woodridge.

Sample Preparation:

1. Jaw Crusher

Procedure

This device crushes the sample to a nominal 10mm by the repeated compression of two steel plates. When used it is let to run continuously and is cleaned with compressed air in-between each sample.

Comments/Problems

This method of cleaning appears to be sufficient in eliminating contamination as the jaw crusher’s workings are fairly simple.

2. Roller Crusher

Procedure

Samples are then inserted into the roller crusher. It consists of two closely spaced rollers through which the sample is fed through and crushed. The sample is crushed to a nominal 2-3 mm, however tabular chips have been seen as long as 20mm. The machine is let to run continuously and is cleaned in-between each sample using compressed air.

Comments/Problems

Though apparently free of loose particulate matter, a very thin but suspicious crust lines the rollers which may possibly affect sample purity.

3. Splitter

Procedure

Crushed sample is then fed into a splitter that splits the sample 50:50. 1kg of sample is then retained for testing while the rest is bagged and returned to the client as coarse rejects.

4. Drying

Procedure

The sample is then bagged and dried in a kiln at 80 to 130 deg. Celsius.

5. Keegor Mill

Procedure

This is the final stage of the grinding/crushing process. The process is similar to the grinding of flour in a stone flour mill.

The 1kg sample is fed into the hub of two oppositely rotating mild-steel plates (somewhat reminiscent of the stone wheels found in a flour grinding mill) and the resultant crushed material is expelled from their rim and collected via a schute and into a pulp reject bag. The controlling factor of discharge is size, so no sample leaves until crushed down to 100 microns. This may mean long grinding times for very hard rock samples.

Of the original 1Kg of material inserted into the Keegor mill, 400g is retained for testing and the rest is discarded. The Keegor mill then is flushed with stream sediments after each sample. An assay of the stream sediments is taken for each load received, and is generally below 10 ppb Au.

Comments/Problems

It is questioned whether a single assay is sufficient to assure the sterility of an entire load of stream sediments, which may be up to a ton, due to the possible heterogeneity of the stream sediments. The lab manager has explained that the sediments are at least somewhat homogenized by the shoveling procedures used to appropriate them.

It should be noted that the mild steel plates (Burnell hardness of 250 to 280) wear away as each sample passes through them and as a result each sample is slightly contaminated with mild steel. Although this steel contamination does not have a role in affecting final assayed Au quantities, it is important to ensure that the grinding plates are of a high purity of mild steel, for case examples have been noted where contamination of other elements such as Sn and Cu have resulted from the use of impure grinding plates, and assay results for these elements have been significantly altered by such impure grinding plates. Harder grinding plates were not recommended by the lab manager due to resultant overheating.

Firing of Sample

6. Mixing of Sample and Flux

Procedure

50g (plus or minus .2 g) of the ground sample is then combined with approximately 150g of flux. The flux contains sodium carbonate, lead oxide, hydrated sodium borate, silica, and flour. This is mixed in a plastic container and then poured into a crucible along with 1 ml of silver nitrate solution. The crucibles are not labeled with their correspondent sample numbers, but are arranged in a 5x10 grid pattern, with number “0” being in the lowest left, and number “49” being in the highest right. Copper sulphate powder is always added in the number “0” crucible, and is also added into the crucible representing the firing of that day. For example, if the sample preparer is preparing samples for the 3rd firing of the day, then he would put copper sulphate powder in crucible number “0” and in crucible number “3.”

1 ml of silver nitrate solution is also added to the crucibles during this time which is necessary for the proper dissolution of gold later in the assaying process. The Ag/Au ratio in the sample must always be above 3:1, and is preferred by the lab manager to be kept at approximately 15:1.

Comments/Problems

The plastic container used for mixing the sample with the flux is not cleaned in-between samples. It is simply banged against a hard surface to knock out whatever loose material can easily come out. Suspicions arise about whether banging is an effective enough method of cleaning the plastic container. A barren flush would be effective in eliminating any chance of contamination.

Also, because there is a margin of error allowance of plus or minus .2g, a .4 percent error potentiality is immediately injected into the assay process. Decreasing the margin of error at this stage in assaying would be a simple way of increasing precision of assay results.

7. Reducing Fusion Furnace

Procedure

Up to 50 crucibles are then placed into the reducing furnace at a time. The flux and sample become liquefied and through that liquid Pb alloys itself with the Au and Ag added to the mix during the fluxing stage. The samples are fired for 45 to 50 minutes if the sample is mostly oxides, and up to 75 to 90 minutes if the samples are complex silicates.

Comments/Problems

It is up to the assay operator’s discretion to determine whether the sample is of a simple oxide or a complex silicate, and to determine firing times. Misidentification of the sample-type may result is insufficient firing times and lead to “bad fusion” (see below) and inaccurate assay results.

8. Pouring/Knocking Out

Procedure

After a given amount of time the crucibles are removed from the reducing furnace and the molten sample is poured into cone-shaped moulds where the molten material may solidify. During this stage each crucible used three times before it is discarded. The PbAuAg alloy settles to the bottom of the mould while still in a liquid state. After the molten sample had solidified, it is removed from the mould and the PbAuAg alloy “button” which is less than 1cc in volume is collected and placed into a smaller crucible.

Comments/Problems

In most samples that are poured, the PbAuAg alloy cannot be seen as it is poured from the crucible into the mould and settles to the bottom of the mould during cooling. However, in some samples the PbAuAg has not alloyed itself together properly and can literally be seen pouring off as chunks into the moulds. This is known as “bad fusion” and can result in inaccurate assay results due to the following circumstances: A. The PbAuAg alloy may be retained in the crucible after pouring and discarded, resulting in loss of Au form the assay procedure and inaccurate assay results; B. The PbAuAg may have not alloyed properly in the molten state and some Au may still be retained in the vitreous substrate.

Both of these circumstances result in a smaller than normal PbAuAg button. It was explained by the lab manager that the assay operator watches for signs of bad fusion and takes note of them, marking “SB” for “small button” in a log, but there seems to be a high potentiality for these signs to be overlooked or ignored, resulting in imprecise assay results.

It should also be noted that in addition to Pb, Au, and Ag, platinum and palladium are also alloyed into the button but are effectively removed during later assay stages.

9. Oxidizing Fusion Furnace

Procedure

PbAuAg alloy buttons are placed in small crucibles and inserted into the oxidizing furnace where the sample is blasted with oxygenated air superheated to 1000 deg. Celcius. During this process the Pb oxidizes and is absorbed into the crucible which becomes discolored, leaving a small pellet of AuAg. Pellets observed were up to 2 mm, but larger ones may occur, depending on Au quantities.

During this stage, the copper sulphate added to the sample during the fluxing stage causes the crucible to become a blue color. This ensures that throughout the process the orientation of the crucibles has not changed and the samples are in their correct order. The copper sulphate stained crucibles used during this stage are used once and discarded. All other crucibles used during this stage are used twice and discarded.

Comments/Problems

It is not possible that these AuAg pellets can roll out of their cucibles, as they are slightly adherent to the crucible, and must be deliberately removed with pliers.

The greatest potential problem during this stage is confusion of samples, since they are not numbered individually. There are several techniques used to alleviate this problem, such as specially-shaped crucible holders which can be racked only in one direction, and the already mentioned copper sulphate staining technique. No doubt such careful measures are taken due to learning, the hard way.

Laboratory Preparation of Sample/Analysis

10. Dissolution, Precipitation

Procedure

The AuAg (still with traces of platinum and palladium) are removed from the crucibles and placed into glass test tubes. The number “0” sample is labeled with “Cu” along with the sample number that represents the firing of the day.

.5ml of concentrated nitric acid (36%) is then added to each test tube. The samples are placed onto a hot plate to speed the dissolution of Ag and Au. This dissolution takes approximately 15 minutes.

Hydrochloric acid is then added, causing the Ag to precipitate out of solution and settle to the bottom of the test tube, leaving a solution of dissolved Au.

Comments/Problems

The test tubes are discarded after one use. The acids are added to the sample using an automatic dispenser that is claimed to have les that 1% error in accuracy.

11. Au Concentration Testing/AAS

Procedure

Within 15 to 20 minutes the silver nitrate is completely precipitated and the sample is ready to be analyzed by the AAS. Without centrifuging or decanting, a small tube is inserted into the sample test tube and a small amount of the Au solution is carefully drawn out so as to not disrupt the silver nitrate precipitate. The Au solution is then sprayed into the air-acetylene flame. A spectrograph is used to measure the strength and spectrum of light that the flame produces. When the Au solution is sprayed into the flame, it produces an absorption spectrum which is detected by the spectrograph and measured for strength. Based on the strength of the absorption spectrum measured, the concentration of the Au solution and thus the mass of Au present in the sample can be deduced and is presented on a digital display.

Each sample is measured three times, and each measurement is taken over an interval of two seconds.

The AAS must be calibrated before use using several standard Au solutions. The AAS can then only measure Au quantities falling in the range of the standard solutions used to calibrate it. The AAS is not capable of being calibrated above 900 ppb Au, so if the sample solution is too strong for the machine then the sample will have to be diluted and this dilution factor will have to be taken into account when reading the AAS display.

Pure water is sprayed into the AAS every 10 samples. The data is then hand written and entered into a computer. Throughout the entire assaying process, there is an error of 15% in accuracy.

Comments/Problems

Care must obviously be taken in operating this machine. Errors may occur in calibrating, i.e. using wrong standard solutions or incorrect typing of concentrations into machine. Errors may occur in hand writing of data from the digital display, and during data entry into the computer.