## Slide 1: Data report
Report to | Stephen Pearson
Reported by | Dr Renjie Zhou
RIF Director | Professor Jianxin Zhao
RIF Manager | Dr Nicole Leonard
Report date | 18 Nov 2024

## Slide 2
Summary of samples and project timeline
Methods
QA/QC
Plots of results
What’s in this document?

## Slide 3
Summary of samples and project timeline
Materials/Samples: One metamorphic rock
  24GPDD001- 632.14m-632.92m (Lab code: 24GPDD001_632)
Analysis: Detrital zircon U-Pb geochronology by laser ablation ICP-MS with at least 100 grains dated from each sample
Project timeline
  Initial discussion: late Sept 2024
  Rock samples received at UQ: 26 Sept 2024
  Zircon mineral separates returned to UQ: 30 Oct 2024
  Laser ablation ICP-MS analysis: 16 Nov 2024
  Data reporting: 20 Nov 2024

## Slide 4
Methods
Laboratory, materials, instrument
Lab name | Radiogenic Isotope Facility (RIF), The University of Queensland (UQ)
Materials | Detrital zircons
Sample preparation | Conventional mineral separation conducted by Geotrack International Pty Ltd One-inch resin round mounts, ~0.3 μm polish to finish
Laser ablation system | ASI RESOlution 193nm excimer laser ablation system
ICP-MS system | Thermo Scientific™ iCAP™ RQ Single Quadrupole (SQ) Inductively Coupled Plasma Mass Spectrometry (ICP-MS) system
Laser ablation operating conditions
Spot size | 30 μm round spot
Repetition rate | 6 Hz
On-sample fluence | 2.9 J/cm2
Carrier gas | 320 ml/min He gas mixed with 5 ml/min high-purity N2
ICP-MS operating conditions
RF Power | 1550 W
Nebuliser gas | Ar gas flow at 0.94 l/min
Sampling depth | 4.0 mm
Mass fractionation | Monitored with NIST612 glass: 232Th/238U = ~1.0, 206Pb/236U = ~0.23

## Slide 5
QA/QC
Primary reference material: 91500 zircon (1065 Ma)
  Zircon reference material 91500 is one of the best characterised geochemical reference materials for microanalysis. It is a single crystal, natural zircon reference material accepted as such by both the secondary ion mass spectrometry (SIMS) and the laser ablation inductively coupled plasma mass spectrometry (LA-ICP-MS) communities.
  Citation: Wiedenbeck, M.A.P.C., Alle, P., Corfu, F.Y., Griffin, W.L., Meier, M., Oberli, F.V., Quadt, A.V., Roddick, J.C. and Spiegel, W., 1995. Three natural zircon standards for U‐Th‐Pb, Lu‐Hf, trace element and REE analyses. Geostandards newsletter, 19(1), pp.1-23.
Secondary reference material: TEMORA2 zircon (417 Ma)
  TEMORA2 zircons are natural zircon grains from the Middledale Gabbroic Diorite within the Palaeozoic Lachlan Orogen of Eastern Australia. It is a major international reference material mainly distributed by Geoscience Australia.
  Citation: Black, L.P., Kamo, S.L., Allen, C.M., Davis, D.W., Aleinikoff, J.N., Valley, J.W., Mundil, R., Campbell, I.H., Korsch, R.J., Williams, I.S. and Foudoulis, C., 2004. Improved 206Pb/238U microprobe geochronology by the monitoring of a trace-element-related matrix effect; SHRIMP, ID–TIMS, ELA–ICP–MS and oxygen isotope documentation for a series of zircon standards. Chemical Geology, 205(1-2), pp.115-140.

## Slide 6
QA/QC
Results from primary and secondary reference materials are consistent with their literature values with both excellent precision and accuracy.
Primary reference material 91500 zircon
Concordant age: 1062.4 ± 1.2 Ma (2σ level error bar)
Secondary reference material TEMORA2 zircon
Concordant age: 418.2 ± 1.5 Ma (2σ level error bar)
n = 50
n = 112

## Slide 7
Plots of results
  24GPDD001- 632.14m-632.92m (Lab code: 24GPDD001_632)
  150 spots were measured.
  Overview of dates:
All data
(n = 150)
    A cluster of near concordant dates with a significant trend pointing to the origin.
    This observation is consistent with the lithology (metamorphic rock).
    The disperse of dates (and a few grains sitting over 2000 Ma) implies that the metamorphic events did not fully reset zircon U-Pb system.
    The lower intercept age sits at the Paleozoic but this information is not robust enough to be interpreted to a geological event.

## Slide 8
Plots of results
206/207Pb ages
(n=74)
  Filtering
    Given the large number of analyses, a conservative concordance filter (-2% to 5%) is applied.
    74 concordant analyses yield a peak at around ~1830 to 1850 Ma.
Weighted mean of
206/207Pb ages:
1838 ± 5 Ma